## Slide 1
Zygo v3
PyQt6
Zygo Steps for acquisition and analysis
Get Images
Images are intensity measurements of interferences. A set of 5 images is necessary to be demodulated by the Hariharan phase demodulation algorithm.
They can be obtained by Acquisition process (required Hardware) or by opening a MAT file.
Data can be stored in MAT file, containing "Images" (set of 5 arrays in 2 dimensions), "Masks" objects (array(s) in 2 dimensions - same size as images) and "Masks_Type" list.
Acquisition
A set of 5 images are acquired through an IDS camera after moving a piezo actuator at 5 different position : 0, λ/4, λ /2, 3 λ/4, λ - where λ is the wavelength of the source. Optical system must be aligned !
Process Phase
Images are masked and then used for processing phase (wavefront) by the Hariharan phase demodulation algorithm.
A first step calculates the wrapped phase. A second step calculates the unwrapped phase.
Add Masks
Masks allow to select a part of the images. A set of different masks can be added (circular, rectangular, polygonal).
Analyses
Different kinds of analyses can be performed on the wavefront :
Simple view of the wavefront (2D or 3D)
Correction of the wavefront (based on Zernike coefficients)
PSF, MTF, circled energy calculations and displays

## Slide 2
Zygo v3
PyQt6
MVC pattern
Models
DataSetModel
  ImagesModel
  MasksModel
  PhaseModel
Views
MainView
MainMenu
Surface3D
Surface2D
Surface3Dx2
ImageView
CoefficientsTable
MasksManagerView
MaskZoomView
AcquisitionZoomView
HtmlView
ProgressionBar
…
Controllers
ModesManager
ImagesController
ZygoApp
+ data_set
+ main_widget
+ main_menu
+ mode_manager
Modes : Acquisition, Images, Masks, Analysis, Aberrations, Simulation, Correction, Help, Options
Controllers manage submenu

## Slide 3
Zygo v3
PyQt6
Modes
Acquisition
Images
Masks
Analysis
Aberrations
Correction
Simulation
Help
Options
Camera, Piezo, Zoom for acquisition, Simple Acquisition, Multiple Acquisition
Open a set of images, Diplay a set of images
Open a mask, Display masks, Manage/Create new masks (C, R, P)
Process Wrapped and Unwrapped phase, PV and RMS calculations, Tilt Correction (via Zernike)
Process Zernike/Seidel coefficients, Strehl ratio, PSF, MTF, Circled energy
Display help pages (Alignment, Zygo, aberrations…)
Change langage, change lambda
Create simulated wavefront from Zernike coefficient
???

## Slide 4
Zygo v3
PyQt6
Top_left
Top_right
Bot_right
MainMenu
Views / MainWidget
Submenu
Options1
Options2
Bot_left

## Slide 5
Zygo v3
PyQt6
Views / MainMenu
options_list
Allows to unactive buttons in the main menu depending on the model (hardware, data, analysis)
+ nocam : No camera is connected (acquisition is not possible)
+ nopiezo : No piezo is connected (acquisition is not possible)

## Slide 6
Zygo v3
PyQt6
Acquisition
Images
Masques
Analyse Simple
Aberrations
Simulation
Options
Piezo

## Slide 7
Zygo v3
PyQt6
Aide
Ouverture Set
Gestion Set
Sauvegarde
1
2
5
??
Set
Masques
1
Acquisition
Images
Masques
Analyse Simple
Aberrations
Simulation
Options
Piezo

## Slide 8
Zygo v3
PyQt6
Aide
Circulaire
Rectangulaire
Polygonal
Liste des masques
1 [SHOW] [INVERT] [DEL]
G [SHOW] [INVERT] [DEL]
Acquisition
Images
Masques
Analyse Simple
Aberrations
Simulation
Options
Piezo

## Slide 9
Zygo v3
PyQt6
Non déroulée
Phase Déroulée
Wedge Factor […]
PV RMS
1 5,4 1,3
Acquisition
Images
Masques
Analyse Simple
Aberrations
Simulation
Options
Piezo

## Slide 10
Zygo v3
PyQt6
Non déroulée
Phase Déroulée
Wedge Factor […]
PV RMS
1 5,4 1,3
Correction
[] Tilt
PV RMS
1 5,4 1,3
Acquisition
Images
Masques
Analyse Simple
Aberrations
Simulation
Options
Piezo

## Slide 11
Zygo v3
PyQt6
Coefficients (table)
Zernike
Seidel
Non corrigée / PV = 5,4 - RMS = 1,3
Zernike (0 -> 36)
Acquisition
Images
Masques
Analyse Simple
Aberrations
Simulation
Options
Piezo

## Slide 12
Zygo v3
PyQt6
Caméra
Images
Masques
Acquisition
Analyse Simple
Aberrations
Correction
Analyse Poussée
Options
Piezo
Correction
[] Tilt
Non corrigée / PV = 5,4 - RMS = 1,3
Corrigée/ PV = 5,4 - RMS = 1,3
Zernike (0 -> 36)
[] Defocus
[] Astigmatisme
[] 3 [] 5 [] 7 [] 9
[] Coma
[] 3 [] 5 [] 7 [] 9
[] Ab. Sphériques
[] 3 [] 5 [] 7 [] 9 [] 11
[] Trefoil
[] 5 [] 7 [] 9

## Slide 13
Zygo v0
OLD

## Slide 14
Zygo v0
OLD

## Slide 15
Zygo v2
PyQt6
Caméra
Images
Masques
Acquisition
Analyse Simple
Aberrations
Correction
Analyse Poussée
Options
Piezo
Spot Diagram
PSF
FTM

## Slide 16
Zygo v2
PyQt6
Caméra
Images
Masques
Acquisition
Analyse Simple
Aberrations
Correction
Analyse Poussée
Options
Piezo